## Slide 1
ENGINEERING
THE CHANGE
Notes: <number>

## Slide 2
EU MDR (2017/745)
The new EU MDR will replace the EU’s current Medical Device Directive (MDD 93/42/EEC) and the EU’s Directive on Active implantable medical devices (AIMDD 90/385/EEC).
The Medical Device Regulation was officially published on May 5th, 2017 and came into force on May 25th, 2017. Manufacturers of currently approved medical devices had a transition time until May 26th, 2020 (revised to May 26th 2021) to meet the requirements of the MDR. MDD certificates that expire before 5/25/2020 can be extended up to 5/25/2024.
EUMDR Move towards more details…
MDD
60 pages
23 Articles
12 Annexes
EU MDR
175 pages
123 Articles
17 Annexes
10 Chapters

## Slide 3
EU MDR (2017/745)
Potential impact on Medical Device Manufacturers

## Slide 4
Major Changes (1/3)
MDD 93/42/EEC
EU MDR
Common Specification (CS)
Manufacturers/operators/conformity assessment bodies can use harmonized standards to demonstrate products/processes comply to EU regulations
Comply with CS OR justification of adopted solutions ensure level of safety and performance.
Risk, General Safety & Performance
Annex I Essential Requirements focused on Safety
-Takes the place of current Essential Requirements
-Adopt continuous iterative risk management system
-Evaluate impact from production phase and PMCF
-Manage risk controls to improve residual risk
Post Market Surveillance
The clinical evaluation and its documentation must be actively updated with data obtained from the post-market surveillance.
PMS system which shall be an integral part of the manufacturers quality management system
Traceability of Devices
Annex I section 13.3. Non electronic labeling and product identification, which are EC State specific.
Unique identification number to provide for traceability throughout the supply chain to end-user/patient,
The European Databank on Medical Devices (Eudamed) is expected to provide information on device data (Publicly & Restricted).
Scope
Covers Medical devices intended by its manufacturer to be used for diagnostic , prevention, monitoring, treatment or alleviation of diseases.
Accessories brought under Medical devices classification.
Extended Scope: Eg., -Contact Lenses / Cleaning, Disinfection or sterilisation of medical devices. Refer Annex XVI for list.
Technical File (Technical Documentation)
Annex VII section 3. Dossier focused on Safety and not on Clinical Efficacy
Requirement of additional details, such as
Reference to previous & similar generation of device
PMS plan, Product V&V details.
Notes: Original MDD 93 / 42/ EEC was Published on 14 June 1993
Technical File:
Benefit Risk Analysis
Results of Preclinical and clinical data
<number>

## Slide 5
(2/3)
Product Classification
Recertification
Clinical Evaluation
…
No “grandfathering” provisions – Under the MDR, all currently approved devices must be recertified
Clinical Evaluation must follow a procedure based on:
1.1.1. either a critical evaluation of the relevant scientific literature currently available relating to the safety, performance, design characteristics and Intended purpose of the device
There is demonstration of equivalence of the device to the device to which the data relates, and
The data adequately demonstrate compliance with the relevant essential requirements.
1.1.2. OR a critical evaluation of the results of all clinical investigations made.
1.1.3. OR a critical evaluation of the combined clinical data provided
Establish and update a clinical evaluation plan
Clinical evaluation can only be based on clinical data of a similar device
Clinical investigation shall address intended purpose, clinical benefits, performance & device safety
Clinical investigations shall be subject to scientific and ethical review.
Post market clinical follow up becomes mandatory and integral to update risk management and clinical evaluation for all medical devices.
Active Implantable devices are covered under Directive 90/385/EEC
Surgical Meshes - Not Covered
No Separate classification for product software (i.e., Software, which drives a device or influences the use of a device, falls automatically in the same class.)
NEWLY ADDED
Surgical meshes” added as class III devices
Product Software is classified as IIa, IIb or III based on intent of use and severity
All devices incorporating or consisting of nanomaterial are in class III
Devices with substances that are intended to be introduced into the human body or applied on skin and that are absorbed by or locally dispersed in the human body are: in class III
MDD 93/42/EEC
EU MDR

## Slide 6
Major Changes (3/3)
Qualified Person
Notified Body (NB) And conformity Assessment
Distribution Chain
No Such Requirement
Identification of qualified person.
Document specific qualifications of this individual relative to the required tasks.
Regulatory representative for all manufacturers.
Notified body shall inform its competent authority about all certificates issued, modified, supplemented, suspended, withdrawn or refused and the other notified bodies within scope of this Directive about certificates suspended, withdrawn or refused and, on request, about certificates issued. The notified body shall also make available, on request, all additional relevant info.
Increased scrutiny from NB in reviews and QMS audits
Increased NB auditing requirements (Unannounced Audits)
NB shall notify the competent authorities, certificates it has granted to devices, the conformity assessment of which has been performed.
No Such Requirement
General obligations of distributors, such as:
Manufacturer has complied with import requirements
Manufacturer has Complied with UDI requirements
Ensure storage and transport is as set out by manufacturer
Also in certain cases obligations of manufacturers apply to distributor, importer
MDD 93/42/EEC
EU MDR
Notes: <number>

## Slide 7
EU MDR (2017/745)

## Slide 8
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
Being certified to ISO 13485 is important, this does not ensure that QMS will be current with the new EU MDR criteria. EU MDR introduces some additional QMS expectations with respect to:
Post-market Surveillance System & Periodic Safety Update Report (PSUR)
  One of the big QMS changes is to post-market surveillance. With the life cycle approach in focus, so there is an emphasis on continuously updating risk documents and Technical file with results of evaluation.
  This expectation is so important for medical device companies to truly manage their devices throughout the entire product lifecycle-from design and development, through design transfer, through post-market surveillance, through product and process changes.
  Post market Surveillance (PMS) plan is mandatory and should be part of Technical documentation.
  PMSR : Manufacturers also required to produce a post-market surveillance report (PMSR). This is intended for low-risk Class I devices and it must summarize the results and conclusions of Post Market Surveillance (PMS) data along with a rationale and description of any corrective actions taken for products on the market. This report becomes part of technical documentation

## Slide 9
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
Post-market Surveillance System & Periodic Safety Update Report (PSUR)
  The Periodic Safety Update Report (PSUR) is addressed in Article 86 and is essentially an extension of the PMSR, but contains information for devices in higher risk categories. It covers classes IIa, IIb, III and implantable.
  It summarizes the results and conclusions from your PMS data. Additional information provides a summary of post market information, vigilance reporting, and current status of these devices on the market in the EU.
  Manufacturer also have to provide a rationale for any corrective actions. This report must be updated annually for most devices, although Class IIa is only required to do so every two years. Additionally, any Class III devices or implantable must submit their PSUR to the EUDAMED database.

## Slide 10
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
INCIDENTS AND FIELD SAFETY CORRECTIVE ACTIONS
  Incidents and field safety corrective actions are addressed in Articles 87 – 89.
  Manufacturers to make sure that they meet requirements of the new regulations. Must ensure that QMS processes outlines new requirements for reporting incidents and field safety corrective actions.

## Slide 11
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
Economic Operators (EOs)
  EU MDR identifies distributors, importers and EU authorized representatives as Economic Operators
  Each with specific responsibilities regarding verification of compliance, cooperation in complaint handling and field safety corrective actions, and cooperating with manufacturers and Competent Authorities in device traceability. (Article 11, 13 & 14)
Distributors and importers may never have had a quality system before. Under this regulation, they will need to develop one.

## Slide 12
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
Unique Device Identification (UDI) and Labelling & EUDAMED.
  UDI is a new requirement under MDR addressed in Article 27 (This is similar to US FDA UDI requirements)
  EUDAMED is also new to EUMDR (Article 33 & 34)
  Manufacturer need to develop processes for UDI and labeling that:
    Defines storage of the UDI at each level (electronically is preferred).
    Details the process for importing UDI information into EUDAMED.
  Manufacturer also required to keep an up-to-date list of all UDIs.
Note: EUDAMED are deferred requirements. At the time, no one knows how EUDAMED is going to work or what it will look like.
UDI Implementation dates:
Class III devices 26 May 2021.
Class IIa & class IIb devices 26 May 2023.
Class I devices Article 27(4) shall apply from 26 May 2025

## Slide 13
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
Person responsible for Regulatory Compliance (Article 15).
  The EU MDR outlines the requirements of the “person responsible for regulatory compliance.”
  They must have a relevant degree or sufficient work experience to understand and meet the requirements of the standard.
  It also outlines that if you are a small company, you don’t have to have this person on-staff, but you should consult with a company that has the regulatory compliance expertise.

## Slide 14
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
Document Storage & Retention.
  The policy for document and storage retention it has been updated.
  Retain the documentation for devices for 10 years.
  In the case of implantable devices, the period shall be at least 15 years.

## Slide 15
EU MDR (2017/745)
QMS IMPACT FOR EU MDR
Implantable Devices
  If the manufacturer is into implantable devices, you will find that almost all of them are now Class III (Ref, Annex VIII)
  Patient implant cards (Article 18) are required for all implantable medical devices which must be created for each device and reviewed under design control and risk.
  Summary of Safety & Clinical Performance, SSCP (Article 32) for implants and class III devices.
  Additionally
    There are new requirements for evaluation of materials for CMR & EDS (Carcinogenic, Mutagenic, Toxic to Reproduction and Endocrine disturbing substances and also Phthalates. ( Annex I, 10.4)
    Requirements on Harmonized standards : Document how you will comply with and review the harmonized standards.
    Along with CER (Clinical evaluation report) there is an additional Requirement of Clinical Evaluation plan (CEP) and also Post Market clinical follow up (PMCF) and periodic update of CER based PMS & PMCF data.

## Slide 16
Thank you